## Slide 1: Soft X-Ray Spectrometers for Laser Driven Backlighter Experiments
Proposal Presentation by Carlos Butler
Notes: Here we see the experimental chamber for the experiment in may of this year, with spectrometers visible.

## Slide 2
Motivation
Spectrometer Design
Mechanical Design
Outlook

## Slide 3: Warm Dense Matter (WDM)
State between condensed matter and weakly coupled plasma
Relevant for:
  Astrophysics
  Inertial confinement fusion
Testing of current theory
Experimental challenges:
  Extreme conditions
  Short lifetimes (
  Opaque in optical range
Riley D., Plasma Phys. Control. Fusion 60, 2018
Notes: Temperatures from 1 – 100 eV, pressures > 100 GPa and above solid density
Say whole graph is for WDM
Breaks common approximations in theory, since strong coupling and electron degeneracy
Due to experimental challenges, specialized diagnostic techniques required.

## Slide 4: Warm Dense Matter (WDM)
Ion Beam
State between condensed matter and weakly coupled plasma
Relevant for:
  Astrophysics
  Inertial confinement fusion
Testing of current theory
Experimental challenges:
  Extreme conditions
  Short lifetimes (
  Opaque in optical range
Sample
Backlighter
target
Intense Laser
WDM
Plasma
X-rays
Setup
Notes: As opaque in optical range, use x-rays
Unique chance to combine heavy ion beam and PW laser at GSI and FAIR
One of the methods using x-rays for diagnostics is the following, which will be used in our experiments

## Slide 5: X-Ray Absorption Fine-Structure (XAFS)
Studies structures around edge (rapid rise) in absorption spectrum
Structures caused by quantum scattering
X-rays
Sample
Transmitted x-rays
Spectrometer
Source
X-ray
Atoms in lattice
Photoelectron wavefunction
Back-scattered wavefunction
For WDM, can extract:
  Density
  Temperature
  Ionic order
  Ionization
Notes: A mature method used in numerous fields (biology, solid-state physics, since 1988 WDM research
Edge is jump in absorption at ionization energy of a shell
Process in diagram happens in sample (point out)
XAFS: 1) x-ray photoionizes electron from absorbing atom
2) photoelectron quantum mechanically scatters on neighboring electron
3) part of wavefunction reflected back to absorbing atom
4) Backscattered wavefunction modulates electronic state at first atom
5) absorption coefficient dependent on electronic state  modulates mu as well  structures dependent on lattice and structure
IMPORTANT: all happens in a single coherent quantum state at absorbing atom  all absorption events effected!
Next will see example of an absorption spectrum

## Slide 6: XAFS Example
X-ray Absorption Near Edge Structure (XANES)
Extended X-ray Absorption Fine-Structure (EXAFS)
For example:
Amplitude of oscillations  temperature
Edge
High T Low T
Photon Energy
Absorption Coefficient

## Slide 7: Planned Experiments
SIS-8 (GSI)
SIS-100 (FAIR)
End goal: XAFS experiments on WDM of aluminum
Proposal goal: Design two spectrometers for testing in laser-only experiment
PW Laser (FAIR)
Spectrometer
(absorption)
Backlighter
target
Aluminum
Sample
Spectrometer
(reference)
X-rays
PHELIX (GSI)
Notes: Final experiment will be at FAIR probing WDM with SIS-100
For that need pre-experiments

## Slide 8: Fundamentals of X-Ray Spectrometers
Crystal
X-ray Source
Detector
Goal: Spectrally disperse photons on detector
X-ray spectrometers:
  Crystal as dispersive element (Bragg condition)
  Properties dependent on crystal form and geometry
In proposal: One flat crystal and one spherically bent crystal geometry  test two spectrometers
Notes: Goal is to disperse wavelengths
Can realize scheme in many different ways. Largely dependent on crystal bending

## Slide 9: Dual Unbent Crystal speCtrometer (DUCC)
View on Detector
Laser Beam
Ion beam
(Future Experiment)
Crystal
Crystal
Detector
Backlighter
Target
Sample
Top view:
Side view:
Laser Beam
Crystal
Detector
Backlighter
Target
Sample
Notes: Two geometries used

## Slide 10: Focusing Spectrograph with Spatial Resolution in 1-D (FSSR-1D)
Two planes:
  Dispersive plane (with Rowland Circle)
  Vertical plane (perpendicular to dispersive)
Spectral focusing in dispersive plane
1-D imaging in vertical plane
Advantage: Better resolution as source size negligible
Source
Dispersive
Spherical crystal
Rowland Circle
Detector
Vertical
Dispersive
Source
On detector
Vertical

## Slide 11: Comparison
DUCC
FSSR
Flat crystal  basic geometry
ADP crystal  excellent intrinsic resolution (due to crystal)
Spherically bent crystal  high intensity and resolution (due to geometry)
Mica crystal  common in literature
Dual channel  only one detector
Small E-range  cover XANES (near edge) at high resolution
1-D imaging  low background
Larger E-range  cover XANES and EXAFS at lower resolution

## Slide 12: Components and Requirements
Main Components | Purpose
Greateyes x-ray camera | Record spectra
Aluminum housing | Structure and shielding
Carbon filters | Insulation of camera chip
Mylar (plastic) filters | Shielding of crystals
Alignment apparatus | Alignment
Requirements:
Implementation of design ( mm precision)
Shielding of sensitive elements
Alignment in chamber
Robustness

## Slide 13: DUCC
Top
Scheme
Front/Side
Front
Back
Camera
Tilt
Blast Shields
Pointer Holder
Carbon Filter Holder

## Slide 14: FSSR
Scheme
Top
Front
Back
Inside Snout
Blast Shield
Crystal
Main Shielding
Snout
Stages for Alignment
Carbon Filter Holder

## Slide 15: Summary
Goal: Design soft x-ray spectrometers for XAFS of Al samples using laser driven backlighters
Design: Two spectrometers
  DUCC: Flat crystal geometry with two channels, FSSR: Spherical crystal geometry with 1-D imaging

## Slide 16: Outlook
Both spectrometers tested in May beamtime (laser-only experiment)
Next steps:
  Develop generalized code for spectra processing
  Assess performance using spectra
  Use results to inform spectrometer design for combined (laser and ion beam) experiment in 2024
Thanks for your attention!

## Slide 17: Considerations
Energy Range
Sample Size
Resolution
Intensity
Physical Size
X-Ray Absorption Fine-Structure: Al K-edge at 1560 eV
Near edge (XANES): 1540 – 1620 eV
Extended (EXAFS): up to 1800 eV
homogeneous heavy-ion beam heating of sample  small sample  small range projection on sample (<1 mm)
Enough to resolve XAFS 
XANES: 1 eV
EXAFS: 10 eV
High as possible without saturation
Fit in HHT chamber  length < 550 mm

## Slide 18: Sample Size Consideration
Crystal
Source
Sample
Ion beam
(Future Experiment)



## Slide 19
Consideration | Limit | DUCC | FSSR
Energy range | XANES: 1540-1620 eV EXAFS: up to 1800 eV | 1541-1618 eV (XANES) | 1465-1755 eV (XANES+EXAFS)
Sample Size | < 1 mm | 0.51 mm | 0.76 mm
Spectrometer length | < 550 mm | 235.2 mm | 404.3 mm
Resolution | XANES: 1 eV EXAFS: < 10 eV | 0.7 eV | 3.1 eV
Consideration | Limit | DUCC | FSSR
Energy range |  | 1541-1618 eV (XANES) | 1465-1755 eV (XANES+EXAFS)
Sample Size | < 1 mm | 0.51 mm | 0.76 mm
Spectrometer length | < 550 mm | 235.2 mm | 404.3 mm
Resolution |  | 0.7 eV | 3.1 eV

## Slide 20: Specifications
Spec | DUCC | FSSR
Energy Range | 1541-1618 eV (XANES) | 1465-1755 eV (XANES+EXAFS)
Intensity on Detector | Lower | High
Resolution | 0.7 eV | 3.1 eV
Physical Length | 235.2 mm | 404.3 mm
DUCC
FSSR

## Slide 21: Simulation of DUCC

## Slide 22: Simulation of FSSR

## Slide 23: Planned Experiments
WDM
experiment
in APPA Cave
at FAIR (>2025)
HHT Cave with PHELIX and SIS-18 at GSI (2024)
HHT Cave with only PHELIX at GSI (2023)
Preparatory experiments
End goal: XAFS experiments on WDM of aluminum
Proposal goal: Design spectrometers for absorption spectroscopy

## Slide 24: Summary
Goal: Design soft x-ray spectrometers for XAFS of Al samples using laser driven backlighters
Design: Two spectrometers
  DUCC: Flat crystal geometry with two channels, FSSR: Spherical crystal geometry with 1-D imaging
Spec | DUCC | FSSR
Energy Range | 1541-1618 eV (XANES) | 1465-1755 eV (XANES+EXAFS)
Intensity on Detector | Lower | High
Resolution | 0.7 eV | 3.1 eV